## CE04A16
D2DH-100 | QuantiChrom™ Lactate Dehydrogenase Assay Kit
Date Released: | 2006-12-05 |  | Protocol Version: | 150512 |  |  |  |  | Date: | 2024-01-01 |  | (e.g. BD04A29, for 130429)
Kit Components: | Qty | Bottle Type |  | Storage | Labels | Weight (g) | Shelf Life (months): | 15 | Lot#: | CE01A01 |  | (mm/dd/yyyy)
Substrate Buffer | 20 mL | 30mL White Bottle |  | -20˚C | -20˚C | 29.8-30.3 |  | Final | By: | AB
Calibrator | 1.5 mL | 2 mL Amber Tube |  | -20˚C | CAL1 -20˚C | 3.141-3.220 | Kit Lot Size: | 56 | Expiry: | 2025-04-01
Diaphorase | 120 µL | 500 µLTube (Yellow Cap) |  | -20˚C | -20˚C | 1.059-1.075 |  |  | Labels Checked By: | PK
NAD Solution | 1 mL | 2 mL Amber Tube |  | -20˚C | -20˚C | 2.677-2.684 |  |  | Date: | 2024-01-02
MTT Solution | 1.5 mL | 2 mL Amber Tube |  | -20˚C | -20˚C | 3.141-3.220 |  |  | QC Checked By: | PK
[9 empty] | Date: | 2024-01-03
Component Name: | Substrate Buffer |  |  | Amount per Batch/Lot: |  | 1200 mL | [6 empty] | x56
Prescribed Chemical and Amounts | [7 empty] | Actual Chemical Used and Qty
Chemical Name | Vendor | Catalog No. | Storage Location | MW | [Final] wt% (mM) | Amount | Units | Vendor | Catalog No. | Lot number | Amount | Units
1 M Tris, pH 8.2 |  | Prepared Stock | Refridg. M | 121.14 | 2.42 wt% (200) | 240 | mL | BAS | stock | 240325 | 240 | mL
Lithium Lactate | sigma | 440469 - 50g | Cabinet B | 96.01 | 0.48 wt% (50) | 5.75 | g | Sigma | 440469-50g | MKCP2871 | 5.7506 | g
dH2O | [5 empty] | to 1200 | mL |  |  |  | to 1200 | mL
Note: Adjust pH to 8.2 with 6N HCL
Component Name: | MTT Solution |  |  | Amount per Batch/Lot: |  | 3000 mL | [6 empty] | x288 (~1000 mL unpacked, stored in Freezer I (or Freezer U))
Prescribed Chemical and Amounts | [7 empty] | Actual Chemical Used and Qty
Chemical Name | Vendor | Catalog No. | Storage Location | MW | [Final] wt% (mM) | Amount | Units | Vendor | Catalog No. | Lot number | Amount | Units
MTT | Biomatik | A3338 | Ref. M | 414.32 | 0.44 wt% (10.67) | 13.263 | g | Biomatik | A3338 | YK2023111001 | 6.5473 | g
dH2O | [5 empty] | to 3000 | mL |  |  |  | to 1480.954 | mL
Combined with 4x CD09G21 and 6x CC04G11
QC Procedures
QC Calibrator: Add 400 µL Calibrator and 400 µL 200 mM NaOH in 50% Isopropanol (Blank) to separate cuvettes.
Measure OD at 565 nm and enter data in Calibrator QC Table below.
QC Reagent: Prepare working reagent: 700 µL of Substrate Buffer + 56 µL MTT + 32 µL NAD + 4 µL Diaphorase
| Standard Curve: Add 5 µL of 0.3 U/µL LDH (See above) for QC to 595 µL of LDH dilution buffer (See Above) to make 2.5 kU/L LDH. | [9 empty] | Used LDH dilution buffer lot 231116. Diluted 0.3 U/uL LDH to 0.8x.
Dilute standards:
(1) 100 µL 2.5 kU/L, (2) 60 µL 2.5 kU/LL + 40 µL Dilution Buffer, (3) 30 µL 2.5 kU/L + 70 µL Dilution Buffer and (4) 100 µL Dilution Buffer.
Add 10 µL of each standard to separate wells.
Detection: Add 190 µL of working reagent to each well. Tap plate to mix.
Read OD 565nm kinetics for 30 minutes. Use data from 0 min and 25 min.
QC Results
Reagents Used For QC:
Acceptable Results:
Reagent | Lot # |  |  |  | Low | High
Substrate Buffer | CE04B16 |  |  | Calibrator ∆OD | 0.8 | 0.87
Calibrator | CE04E10 |  |  | Water ∆OD | 0 | 0.12
Diaphorase | CE01C23 |  |  | Slope | 0.65 | 0.75
NAD Solution | CD12F29 |  |  | R2 | 0.99
MTT Solution | CE04D16
LDH for QC | 211018 |  | Actual Results:
|  |  |  | Calibrator ∆OD: | 0.841 | TRUE
|  |  |  | Water ∆OD: | -0.0001 | TRUE
|  |  |  | Slope: | 0.69024 | TRUE
|  |  |  | R2: | 0.99111 | TRUE
|  | LDH (kU/L) | OD @ 0 min | OD @ 25 min | ∆OD | ∆∆OD
|  | 2.5 | 0.0557 | 1.8193 | 1.7636 | 1.7637
|  | 1.5 | 0.049 | 0.9734 | 0.9244 | 0.9245
|  | 0.75 | 0.0453 | 0.598 | 0.5527 | 0.5528
|  | 0 | 0.0357 | 0.0356 | -0.0001 | 0
Calibrator QC
| Methylene Blue Extinction Coefficient 565 nm (mM-1cm-1): |  |  |  |  | 23.4 |  | Acceptable Results:
[9 empty] | Low | High
| Dil Factor | OD565 | -Blank | X Dil Factor |  | [MethyleneBlue] (µM) |  | [MB](µM): | 34 | 37
| Blank | 0.063 | [5 empty] | %RSD
| 1 | 0.904 | 0.841 | 0.841 |  | 35.9402 |  |  | Pass/Fail
TRUE
NOTE: For Stability data, please see D2DH_2019 BJ12A11

## CofA
Product Information
D2DH-100
Product Lot#: 45292
Quality Control Results
Components | Volume | Vial Code | Test | Criteria | Results
Substrate Buffer | 20 mL | CE04B16 | LDH Standard Curve & Calibrator ΔOD (in Cuvette) | Slope = 0.65 - 0.75 | 0.69024
Calibrator | 1.5 mL | CE04E10 |  | Linearity R2 > 0.99 | 0.99111
Diaphorase | 120 µL | CE01C23 |  | Calibrator ∆OD = 0.8 - 0.87 | 0.841
NAD Solution | 1 mL | CD12F29 |  |  | Pass
MTT Solution | 1.5 mL | CE04D16
Quality Assurance Statement
The information on this certificate of analysis has been reviewed and accurately reflects the
manufacturing and analysis data for the product lot specified above.
Quality Assurance Manager: FH
#VALUE!
Support
For technical support, please call 1-510-782-9988, or E-mail: info@bioassaysys.com