## CE03A01
DALP-250 | QuantiChrom™ Alkaline Phosphatase Assay Kit | [9 empty] | Today
Date Releaded: | 2005 |  | Protocol Version: | 101008 |  |  |  |  | Date: | 2024-03-01 | 2024-06-18 | (e.g. BD04A29, for 130429)
Kit Components: | Qty | Bottle Type |  | Storage | Labels | Weight (g) | Shelf Life (months): | 12 | Lot#: | CE03A01 |  | (mm/dd/yyyy)
Assay Buffer | 50 mL | 60mLwhite Bottle |  | 4˚C | 4˚C | 62.6-64.9 |  |  | By: | AB
Mg Acetate | 1.5 mL | 2 mL White Tube |  | 4˚C | 4˚C | 2.865-2.944 | Kit Lot Size: | 78 | Expiry: | 2025-03-01
pNPP Liquid | 600 μL | 2 mL Amber Tube |  | -20˚C | -20˚C | 2.268-2.293 |  |  | Labels Checked By: | PK
Calibrator | 10 mL | 15 mL White Bottle |  | 4˚C | CAL2 4˚C | 17.0-17.3 |  |  | Date: | 2024-03-02
[9 empty] | QC Checked By: | PK
See 447CD05A18 | [8 empty] | Date: | 2024-03-03
QC Procedures
QC Calibrator: 200 uL H2O, 200 uL Calibrator. OD405nm H2O = 0.04-0.06, ODcalibrator = 0.43 to 0.46
QC Reagent:
(1). Preapre Substrate: mix for each assay 4 uL pNPP + 10 uL MgAc2 + 400 uL Buffer (10mM).
(2). Test Reaction: run duplicate 5 uL bovine serum (from -80C freezer) + 195 uL reconstituted substrate, measure OD405 at 0 and 4 min.
(3). Calibrator: 200 uL H2O and 200 uL tartrazine.
(4). Calculation: DALP Activity = (OD4min - OD0min) x 200/[(ODcalibrator - ODH2O)/5/4] x 35.3 = 353 x (OD4min - OD0min)/(ODcalibrator - ODH2O).
QC 0.25% Tartraziner:
Make three 50X dilutions of Calibrator in dH2O
Add 400 uL of each dilution to separate low volume cuvette (min vol 300 uL) and also add 400 uL dH2O to a separate cuvette to serve as the blank.
Read absorbance of each cuvette at 410 nm on the SpectraMax.
Record Absorbance values in appropriate cell of Standard QC table.
If needed adjust 0.025% Tasrtrazine with either more Tartrazine (low values) or dH2O (high values) to achieve correct absorbance at 410 nm.
QC Results
[5 empty] | Low | High
Reagent | Lot # |  | 200 µL Calibrator ∆OD: |  | 0.37 | 0.42
Assay Buffer | CD05B18
Mg Acetate | CD05C18
pNPP Liquid | CD05D18 |  |  |  | Low | High
Calibrator | CD05E18 | Bovine Serum ∆OD: |  |  | 0.035 | 0.06
Bovine Serum | 1218838
[5 empty] | Low | High
|  | ALP Activity (U/L): |  |  | 30 | 57
|  | OD 5 min | ∆OD @ 405nm |  | Calibrator Pass/Fail
| Calibrator | 0.4493 | 0.4096 |  | TRUE
| H2O | 0.0397 | 0
|  | OD 0 min | OD 4 min | ∆OD @ 405nm |  | Avg. B.S. ∆OD |  |  | Bovine Serum Pass/Fail
Bovine Serum |  | 0.3442 | 0.3834 | 0.0392 |  | 0.04645 |  |  | TRUE
Bovine Serum |  | 0.3291 | 0.3828 | 0.0537
200 uL H2O |  | N/A | 0.04 | 0 |  | ALP Activity |  |  | ALP Activity Pass/Fail
200 uL Calibrator |  | N/A | 0.4437 | 0.4037 |  | 40.6164 |  |  | TRUE
ΔOD
0.25% Tartrazine QC
| Tartrazine Extinction Coefficient 410 nm (mM-1cm-1): |  |  |  |  | 22.66 |  | Acceptable Results:
[9 empty] | Low | High
| Dil Factor | OD410 | -Blank | X Dil Factor |  | [Tartrazine] (mM) |  | [Tartrazine](mM): | 4.5 | 4.9
| Blank | [6 empty] | %RSD | 5
| 50 |  | 0 | 0 |  | 0
| 50 |  | 0 | 0 |  | 0
| 50 |  | 0 | 0 |  | 0
Pass/Fail
[5 empty] | Average | 0 | FALSE
[5 empty] | StDev | 0
[5 empty] | %RSD | #DIV/0! | #DIV/0!

## CD01A03
DALP-250 | QuantiChrom™ Alkaline Phosphatase Assay Kit | [9 empty] | Today
Date Releaded: | 2005 |  | Protocol Version: | 101008 |  |  |  |  | Date: | 2024-01-21 | 2024-06-18 |  | (e.g. BD04A29, for 130429)
Kit Components: | Qty | Bottle Type |  | Storage | Labels | Weight (g) | Shelf Life (months): | 12 | Lot#: | CE01A21 |  |  | (mm/dd/yyyy)
Assay Buffer | 50 mL | 60mLwhite Bottle |  | 4˚C | 4˚C | 62.6-64.9 |  |  | By: | AB
Mg Acetate | 1.5 mL | 2 mL White Tube |  | 4˚C | 4˚C | 2.865-2.944 | Kit Lot Size: | 78 | Expiry: | 2025-01-21
pNPP Liquid | 600 μL | 2 mL Amber Tube |  | -20˚C | -20˚C | 2.268-2.293 |  |  | Labels Checked By: | PK
Calibrator | 10 mL | 15 mL White Bottle |  | 4˚C | CAL2 4˚C | 17.0-17.3 |  |  | Date: | 2024-02-02
[9 empty] | QC Checked By: | PK
[9 empty] | Date: | 2024-02-03
Component Name: | Assay Buffer |  |  |  | Amount per Batch/Lot: | 10 L | [6 empty] | Combined lots CA02B03, CA07B24,
Prescribed Chemical and Amounts | [7 empty] | Actual Chemical Used and Qty |  |  |  |  | BI09B06, CC03B15
Chemical Name | Vendor | Catalog No. | Storage Location | MW | [Final] wt% (mM) | Amount | Units | Vendor | Catalog No. | Lot number | Amount | Units
2-Amino-2-methyl-1-propanol | Sigma | A65182-1L | Cabinet C | 89.14 | 11.6% (1200) | 1160 | mL
dH2O | [5 empty] | to 8.5 | L
6N HCl | BAS | Stock | pH meter | 36.5 |  | to pH 10.5
dH2O | [5 empty] | to 10 | L
Component Name: | Mg Acetate |  |  |  | Amount per Batch/Lot: | 300 mL | [6 empty] | 160
Prescribed Chemical and Amounts | [7 empty] | Actual Chemical Used and Qty
Chemical Name | Vendor | Catalog No. | Storage Location | MW | [Final] wt% (mM) | Amount | Units | Vendor | Catalog No. | Lot number | Amount | Units | Some spilled
Magnesium acetate (MgAc2 4H2O) | Sigma | M2545 | Cabinet B | 214.5 | 4.20% (196) | 12.7 | g | sigma | M2545 | SLCG0244 | 12.76 | g
dH2O | [5 empty] | to 300 | mL |  |  |  | 300 | mL
Note: Pass through sterile filter before packaging. Use 500 mL vacuum sterile bottle (stored in side room next to the cell culture lab).
Component Name: | pNPP Liquid |  |  |  | Amount per Batch/Lot: | 100 mL | (pack 140X 600 uL tubes for DALP and 55X 280 uL tubes for POPN)
Prescribed Chemical and Amounts | [7 empty] | Actual Chemical Used and Qty
Chemical Name | Vendor | Catalog No. | Storage Location | MW | [Final] wt% (mM) | Amount | Units | Vendor | Catalog No. | Lot number | Amount | Units
p-nitrophenol phosphate (4-nitrophenol phosphate) | Santa Cruz | SC-3720A | Refrig. M | 371.14 | 37.11% (1000) | 37.11 | g
dH2O | BAS |  |  |  |  | 20 | mL
*Stir with heat applied (Don't overheat).
phenol | Sigma | 185450-100G | Cabinet B | 94.1 | 53.33% (5667) | 53.33 | g
*Keep stirring with heat applied (Don't overheat). Wait till phenol all melt, adjust pH as below.
Add ethanolamine dropwise (50uL each time) and check pH on pH paper to about 8.0 to 8.5.
Ethanolamine | Alfa Aesar -or- Sigma | A11697 -or- 4111000 | Cabinet C | 61.8 -or- 61.08 | 0.17% (27.19) | about 200 | µL
dH2O | BAS |  |  |  |  | to final 100 | mL
must weigh tube after fill
Component Name: | Calibrator |  |  |  | Amount per Batch/Lot: | 5000 mL
Prescribed Chemical and Amounts | [7 empty] | Actual Chemical Used and Qty
Chemical Name | Vendor | Catalog No. | Storage Location | MW | [Final] wt% (mM) | Amount | Units | Vendor | Catalog No. | Lot number | Amount | Units | used cal from bottle labeled "DALP calibrator"
dH2O | BAS |  |  |  |  | 5000 | mL
0.25% Tartrazine (see Below) | BAS | Stock | Refrig. M |  | 0.0022% | 44 | mL
Component Name: | 0.25% Tartrazine |  |  |  | Amount per Batch/Lot: | 1000 mL
Prescribed Chemical and Amounts | [7 empty] | Actual Chemical Used and Qty
Chemical Name | Vendor | Catalog No. | Storage Location | MW | [Final] wt% (mM) | Amount | Units | Vendor | Catalog No. | Lot number | Amount | Units
Tartrazine | EM | TX0020-3 | Cabinet N | 534.4 | 4.68 (0.25wt%) | 1.25 | g
Acetic Acid | Rocky M | A1053-1gal/cs | Cabinet C | 60.05 | 0.25v% | 1.25 | mL
dH2O | BAS |  |  |  |  | to 500 | mL
See QC Below for 0.25% Tartrazine.
Changes/Notes:
101018RZ: Changed source for p-nitrophenol phosphate used in pNPP Liquid from Hong Kun to Alfa Aesar.
140828RZ: Added QC for 0.25% Tartrazine.
QC Procedures
QC Calibrator: 200 uL H2O, 200 uL Calibrator. OD405nm H2O = 0.04-0.06, ODcalibrator = 0.43 to 0.46
QC Reagent:
(1). Preapre Substrate: mix for each assay 4 uL pNPP + 10 uL MgAc2 + 400 uL Buffer (10mM).
(2). Test Reaction: run duplicate 5 uL bovine serum (from -80C freezer) + 195 uL reconstituted substrate, measure OD405 at 0 and 4 min.
(3). Calibrator: 200 uL H2O and 200 uL tartrazine.
(4). Calculation: DALP Activity = (OD4min - OD0min) x 200/[(ODcalibrator - ODH2O)/5/4] x 35.3 = 353 x (OD4min - OD0min)/(ODcalibrator - ODH2O).
QC 0.25% Tartraziner:
Make three 50X dilutions of Calibrator in dH2O
Add 400 uL of each dilution to separate low volume cuvette (min vol 300 uL) and also add 400 uL dH2O to a separate cuvette to serve as the blank.
Read absorbance of each cuvette at 410 nm on the SpectraMax.
Record Absorbance values in appropriate cell of Standard QC table.
If needed adjust 0.025% Tasrtrazine with either more Tartrazine (low values) or dH2O (high values) to achieve correct absorbance at 410 nm.
QC Results
[5 empty] | Low | High
Reagent | Lot # |  | 200 µL Calibrator ∆OD: |  | 0.37 | 0.42
Assay Buffer | CD01B03
Mg Acetate | CD01C03
pNPP Liquid | CD01D03 | *re-lotted from CC11D14* |  |  | Low | High
Calibrator | CD01E03 |  |  |  | 0.035 | 0.06
Bovine Serum | 1218838
[5 empty] | Low | High
|  | ALP Activity (U/L): |  |  | 30 | 57
|  | OD 5 min | ∆OD @ 405nm |  | Calibrator Pass/Fail
| Calibrator | 0.4473 | 0.4053 |  | TRUE
| H2O | 0.042 | 0
|  | OD 0 min | OD 4 min | ∆OD @ 405nm |  | Avg. B.S. ∆OD |  |  | Bovine Serum Pass/Fail
Bovine Serum |  | 0.3895 | 0.4337 | 0.0442 |  | 0.04195 |  |  | TRUE
Bovine Serum |  | 0.3894 | 0.4291 | 0.0397
200 uL H2O |  | N/A | 0.042 | 0 |  | ALP Activity |  |  | ALP Activity Pass/Fail
200 uL Calibrator |  | N/A | 0.4473 | 0.4053 |  | 36.5368 |  |  | TRUE
ΔOD
0.25% Tartrazine QC
| Tartrazine Extinction Coefficient 410 nm (mM-1cm-1): |  |  |  |  | 22.66 |  | Acceptable Results:
[9 empty] | Low | High
| Dil Factor | OD410 | -Blank | X Dil Factor |  | [Tartrazine] (mM) |  | [Tartrazine](mM): | 4.5 | 4.9
| Blank | [6 empty] | %RSD | 5
| 50 |  | 0 | 0 |  | 0
| 50 |  | 0 | 0 |  | 0
| 50 |  | 0 | 0 |  | 0
Pass/Fail
[5 empty] | Average | 0 | FALSE
[5 empty] | StDev | 0
[5 empty] | %RSD | #DIV/0! | #DIV/0!

## CofA
Product Information
DALP-250
Product Lot#: 45461
Quality Control Results
Components | Volume | Vial Code | Test | Criteria | Results
Assay Buffer | 50 mL | CD01B03 | Serum Control & Calibrator ΔOD | Serum ΔOD = 0.035 - 0.06 | 0.04195
Mg Acetate | 1.5 mL | CD01C03 |  | Calibrator ΔOD = 0.37 - 0.42 | 0.4053
pNPP Liquid | 600 μL | CD01D03 |  |  | Pass
Calibrator | 10 mL | CD01E03
Quality Assurance Statement
The information on this certificate of analysis has been reviewed and accurately reflects the
manufacturing and analysis data for the product lot specified above.
Quality Assurance Manager: FH
Date: 1/0/1900
Support
For technical support, please call 1-510-782-9988, or E-mail: info@bioassaysys.com